## Slide 1: 48-spot Single-molecule detection
WeissLab GM May 1st 2014
Antonino Ingargiola

## Slide 2: Outline
48-spot single-molecule measurements:
  Setup description
  Background characterization (space and time)
  Saturation study
  PSF study

## Slide 3: Current 48-spot setup
Sample: doubly-labeled dsDNA ATT0550-ATTO647N,
we can only see the acceptor

## Slide 4: 48-pixel Manta Layout
 | 0 | 1 | 2 | 3
0 | 23 | 24 | 25 | 26
1 | 21 | 22 | 27 | 28
2 | 19 | 20 | 29 | 30
3 | 17 | 18 | 31 | 32
4 | 15 | 16 | 33 | 34
5 | 13 | 14 | 35 | 36
6 | 11 | 12 | 37 | 38
7 | 9 | 10 | 39 | 40
8 | 7 | 8 | 41 | 42
9 | 5 | 6 | 43 | 44
10 | 3 | 4 | 45 | 46
11 | 1 | 2 | 47 | 48
Upper section
Center section
Lower section
Pixel diameter 50µm
Pixel pitch 500µm
Geometry 12 x 4

## Slide 5: Sample
Donor ATTO550
Acceptor
ATT0647N
  40 base-pair
  Labels: ATTO647N (acceptor) , ATT0550 (donor)
  NHS Ester C6 linker
  Dual HPLC purification

## Slide 6
Donor ATTO550
Acceptor
ATT0647N
1 Sample -> 17 bp D-A distance
  40 base-pair
  Labels: ATTO647N (acceptor) , ATT0550 (donor)
  NHS Ester C6 linker
  Dual HPLC purification
Sample

## Slide 7: Measurements
['/Data/Antonio/data/manta/2014-04-11\\17d_120p_0',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_1',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_2',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_3',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_4',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_400mW_reall_4',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_500mW_reall_nosteer_5',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_500mW_reall_steer_6',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_750mW_reall_steer_7',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_1000mW_reall_steer_8',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_250mW_reall_steer_9',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_400mW_reall_steer_10',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_600mW_reall_steer_11',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_500mW_reall_steer_12',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_750mW_reall_steer_13',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_1000mW_reall_steer_14',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_1000mW_reall_steer_15',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_250mW_reall_steer_16',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_400mW_reall_steer_17',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_600mW_reall_steer_18',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_500mW_reall_steer_19',
'/Data/Antonio/data/manta/2014-04-11\\17d_200p_2_600mW_reall_steer_20']
No steering pattern
Representative (600mW)
Laser powers:
250 mW
400 mW
500 mW
600 mW
750 mW
1000 mW

## Slide 8: Light modulation … to steer on not to steer?
Unmodulated light
background
Modulated light
signal
No beam steering
Beam steering

## Slide 9: Excitation: 13x5, SPAD pitch
No Steering
Vertical Steering
Greys
Fire
ICA

## Slide 10: Results

## Slide 11: 48-spot single-molecule time-trace

## Slide 12: Background vs time (old meas.)

## Slide 13: Background vs time (old meas.)
Why?
Thermal drifts
Laser
…

## Slide 14
… upgrading some optical components …

## Slide 15: Background vs time (new meas.)

## Slide 16: Background
Mean background 4 kcps @ 600mW
cps: counts per second
kcps: kilo-counts per second
(
)

## Slide 17: Number of bursts
Mean number of bursts in 340 sec.
(Burst size > 20ph. @ 600mW)

## Slide 18: Mean background
Convention:
an interrupted line means a new measurement, but the time gap is not to scale

## Slide 19: Mean background
Convention:
an interrupted line means a new measurement, but the time gap is not to scale
Laser power

## Slide 20: Mean background
-25%
Convention:
an interrupted line means a new measurement, but the time gap is not to scale

## Slide 21: Mean background
No beam steering
-25%
with beam steering

## Slide 22: Burst size
The burst size is constant with the laser power

## Slide 23: Burst size
No beam steering
with beam steering
Burst size is constant between steering and no steering

## Slide 24: Burst-per-second (bps)
Doesn’t change in time we are not “loosing” (to many) molecules

## Slide 25: Remarks
Background does not change with time
Burst size does not change with time
Burst-per-second does not change with time
Steering gives 25% background reduction, in this case
Steering increases the signal-to-background ratio (SBR)

## Slide 26: Saturation using the burst size
Bursts are the “survivals” of a burst search and burst selection
Burst search: to be a burst it must have:
    a minimum photon rate
    the min. rate is chosen background rate (F factor)
    background excitation power
Burst selection: to be a burst it must have:
    A minimum burst size of (let say) 20 photons (after corrections).
SBR
Shot-noise
Basic Idea:
bursts are bigger if the emission increases

## Slide 27: Saturation curve
Q: Does this curve implies we are in saturation or not?

## Slide 28: Saturation curve
Q: Does this curve implies we are in saturation or not?
… seems
quite linear

## Slide 29: Saturation curve
Q: Does this curve implies we are in saturation or not?

## Slide 30: Saturation curve
A: we may be in the strongly saturated regime
must be linear
and
passing through
the origin

## Slide 31: Deep thinking the burst size
Bursts are the survivals of a burst search and burst selection
Burst search: to be a burst it must have:
    a minimum photon rate
    the min. rate is chosen background rate (F factor)
    background excitation power
Burst selection: to be a burst it must have:
    A minimum burst size of (let say) 20 photons (after corrections).
    Increasing burst search threshold decreasing the burst size
    Burst size selection differences (we need to subtract offset)
burst size is not a good proxy for the emission intensity
SBR
Shot-noise

## Slide 32: Saturation curve
must be linear
and
passing through
the origin
A: we may be in the strongly saturated regime

## Slide 33: Saturation curve - offset
must be linear
and
passing through
the origin
A: we may be in the strongly saturated regime

## Slide 34: Max. rate in burst
Assumptions
  A selected burst has more than m photons (m=10)
  We use m photons to compute the rate
Max. rate is
  Independent of burst search threshold
Really?
  Yes, to the first order. But there is still a slight dependence:
    Lower rate thresholds more bursts fused higher mean max rate
( a.k.a. the tightest m consecutive photons in a burst )

## Slide 35: Saturation curve – Max rate
Saturation? Not so much below 500mW
Computational details:
Burst search with F=6, size th. = 20
Joined all measurements at same power
Merged all ch into 1 single ch
Computed per-burst max rate
Computed the mean max rate of each power level

## Slide 36: Saturation curve – Background
Saturation? Not so much below 500mW
Computational details:
Burst search with F=6, size th. = 20
Joined all measurements at same power
Merged all ch into 1 single ch
Computed per-burst max rate
Computed the mean max rate of each power level

## Slide 37: Saturation curve – Background
Saturation? Not so much below 500mW
Computational details:
Burst search with F=6, size th. = 20
Joined all measurements at same power
Merged all ch into 1 single ch
Computed per-burst max rate
Computed the mean max rate of each power level

## Slide 38: Open issues
In principle
Saturation happens at high intensity in the spot “core”
Should not be visible in the background because of the lower intensity
We should have “more” saturation in the brightest spots
Experimentally:
Photon rates and BG have the exact same dependence on power
  This means that the “signal” increases as much as the background, indicating no saturation
Also
Subtracting the background is difficult because the max_rate tend to select the background highest tail
Let look at groups of pixels or single-pixels

## Slide 39: Correlation with BG
Burst Search
Min. Rate Th. = const . scales with: Mean( BGCH )
min_rate = const., th=20
min_rate_250mW = 30 kcps
min_rate_XmW = min_rate_250mW * (bg_X/bg_250mW)
Burst Search
Min. Rate Th. BGCH
F = 6,
Bust size th. = 20

## Slide 40
Min. rate @ 250mW = 20kcps

## Slide 41
Min. rate @ 250mW = 30kcps

## Slide 42
Min. rate @ 250mW = 40kcps

## Slide 43
Min. rate @ 250mW = 50kcps

## Slide 44: Saturation in pixel groups
Burst Search
Min. Rate Th. BGCH
F = 6,
Bust size th. = 20
Burst Search
Min. Rate Th. = const . scales with: Mean( BGCH )
min_rate = const., th=20
min_rate_250mW = 30 kcps
min_rate_XmW = min_rate_250mW * (bg_X/bg_250mW)
CH

## Slide 45: Saturation in pixel groups
Burst Search
Min. Rate Th. = const . scales with PLASER
min_rate = const., th=20
min_rate_250mW = 25 kcps
min_rate_XmW = min_rate_250mW *
( P(X mW) / P(250 mW) )
Burst Search
Min. Rate Th. = const . scales with Mean( BGCH )
min_rate = const., th=20
min_rate_250mW = 30 kcps
min_rate_XmW = min_rate_250mW * (bg_X/bg_250mW)
CH
Linear but the rate is too uniform!!

## Slide 46: Refine the burst search
We used
Constant minimum ph. rate threshold for all the CH
But
Different spot have different intensity so different photon rate distribution
Therefore
We clip different distribution with the same threshold and we take the mean
The mean max_rate is not emission
Idea
Selecting only the top N bursts we clip each CH distribution @ “equivalent” points

## Slide 47: Truncated distributions
Distribution mean
Aim: Calculate the ratio of the means
Example: Poisson distribution

## Slide 48: Truncated distributions
Distribution mean
Truncated distribution mean
Aim: Calculate the ratio of the means
Example: Poisson distribution

## Slide 49: Truncated distributions
Distribution mean
Truncated distribution mean
Aim: Calculate the ratio of the means
(is not exactly 2 because of discretization)
Example: Poisson distribution

## Slide 50: Truncated distributions
Distribution mean
Truncated distribution mean
Aim: Calculate the ratio of the means
Simulation: Erlang distribution
Simulation
Form each distribution:
Draw 3000 samples
Selected the top 1000

## Slide 51: Burst search
Thresholding is equivalent of truncating the distribution
Thresholding is necessary to separate the background (another distribution)
Selecting the top N bursts we sample a fixed “fraction” of the distribution
Applying the concept to the burst max_rate we can sample the emission intensity peak
Once we know the peak we can compute the PSF FWHM using the burst duration

## Slide 52: In-burst max-rate
DOES NOT depends on the burst-search threshold
DOES depends on the spot peak intensity
Rationale
The peak count-rate in each bursts excitation
Sampling the “biggest” N bursts:
  we select trajectories that cross the PSF peak
  analogous trajectories in all the CH, irrespective of the PSF intensity
Average across the bursts to smooth the statistical noise
Algorithm
Burst search with Min. Rate Th. = constant
  Avoid background influence
Select top N bursts
  Bursts with highest rate
Compute mean(max_rate) for each channel
  Average the max_rate of each burst
Result
a measure of the PSF peak intensity

## Slide 53: Per-pixel saturation curve
Burst Search
Min. Rate Th. = constant, scales with PLASER
  min_rate = const.
  min_rate_250mW = 25 kcps
  min_rate_XmW = min_rate_250mW * ( P(X mW) / P(250 mW) )
Select top N bursts (3 bps)
Compute: mean(Max_Rate)

## Slide 54: Per-pixel saturation curve
Burst Search
Min. Rate Th. = constant, scales with PLASER
  min_rate = const.
  min_rate_250mW = 25 kcps
  min_rate_XmW = min_rate_250mW * ( P(X mW) / P(250 mW) )
Select top N bursts (3 bps)
Compute: mean(Max_Rate)
 | 0 | 1 | 2 | 3
0 | 23 | 24 | 25 | 26
1 | 21 | 22 | 27 | 28
2 | 19 | 20 | 29 | 30
3 | 17 | 18 | 31 | 32
4 | 15 | 16 | 33 | 34
5 | 13 | 14 | 35 | 36
6 | 11 | 12 | 37 | 38
7 | 9 | 10 | 39 | 40
8 | 7 | 8 | 41 | 42
9 | 5 | 6 | 43 | 44
10 | 3 | 4 | 45 | 46
11 | 1 | 2 | 47 | 48
Upper section
Center section
Lower section

## Slide 55
Saturation curve – pixels groups

## Slide 56: PSF characterization
We can estimate the PSF peak intensity
(mean of top N max_rate)
We can set a burst-search threshold @ 50% of peak
The burst duration becomes to the PSF FWHM

## Slide 57: 48-pixel Manta Layout
 | 0 | 1 | 2 | 3
0 | 23 | 24 | 25 | 26
1 | 21 | 22 | 27 | 28
2 | 19 | 20 | 29 | 30
3 | 17 | 18 | 31 | 32
4 | 15 | 16 | 33 | 34
5 | 13 | 14 | 35 | 36
6 | 11 | 12 | 37 | 38
7 | 9 | 10 | 39 | 40
8 | 7 | 8 | 41 | 42
9 | 5 | 6 | 43 | 44
10 | 3 | 4 | 45 | 46
11 | 1 | 2 | 47 | 48
Upper section
Center section
Lower section

## Slide 58: Max rate distribution

## Slide 59: Max rate distribution: top 800 bursts

## Slide 60: Number of bursts
The # bursts becomes independent from the spot intensity

## Slide 61: PSF Peak intensity

## Slide 62: PSF width

## Slide 63: PSF width
anticorrelation

## Slide 64: Burst size
Open question
Why is the burst size so constant?

## Slide 65: Thank you!

## Slide 66: Burst searches comparison
Burst Search
Th BG
F=6
Size Th=20
Burst Search
Th fixed
Min rate @ 250mW = 25kcps
Min rate scales with laser power
Size Th=20